# Mobile Monitor for Critical Care

AI Driven Diagnostic Support

Project Description

Mobile Monitoring Software Solution (MMSS)

Date: 2026-03-02

# 1. Project Overview

## 1.1 Purpose

The purpose of the MMSS project is to develop medical device software that transforms an existing portable patient monitor from a passive vital signs display into an active clinical decision-support tool. By integrating an AI-driven diagnostic engine into the monitoring software, MMSS enables clinicians — particularly in pre-hospital and mobile care settings — to receive ranked, real-time diagnostic candidates alongside raw vital signs data, converging to a predictive diagnosis within two minutes of patient connection.

The software acquires data from a defined set of non-invasive measurement devices through standardised interfaces, processes it through an AI analysis engine, presents vital signs and diagnostic output on a connected monitor display, generates clinical alarms for abnormal conditions and sensor faults, and optionally shares data with hospital information systems. The project delivers the complete MMSS software stack — from device acquisition through AI processing to user presentation and hospital integration — as a regulated medical device software product compliant with IEC 62304.

# 2. Scope

This project covers software development only. No hardware development, hardware modification, or hardware procurement is in scope. All physical components — the host CPU platform, measurement devices (ECG monitor, pulse oximeter, blood pressure monitor, thermal probe, capnometer, EEG monitor), monitor display, and physical enclosure — are treated as existing, fixed elements. MMSS interfaces with these components through defined interface control documents; their internal design and physical characteristics are outside the scope of this project.

In scope:

Application software only — all five software items: AC, DAC, DEC, DPREC, DPROC

Interfacing with the external Open Evidence AI library

The AI model is used as a commercially validated off-the-shelf component

Drivers and polling logic for the six measurement device types

Vital signs display, alarm presentation, and diagnostic candidate rendering

Hospital information system integration for second opinion

Out of scope:

Hardware design, selection, or modification of any physical component

Device behavior is fixed and accessed through published interfaces

Physical enclosure, portability, weight, or battery design

# 3. Intended Use and Classification

## 3.1 Intended Use

The Mobile Monitor for critical care is intended to provide a comprehensive, real-time monitoring solution for critically ill patients. Its primary purpose is to ensure continuous tracking of key vital signs and offer diagnostic support through advanced technologies such as Artificial Intelligence (AI). The Mobile Monitor is portable, user-friendly, and non-invasive to provide maximum efficiency and comfort in various critical care environments, including intensive care units (ICU), emergency rooms, and mobile medical units. The Smart Mobile Monitor is intended to be used by trained medical professionals.

## 3.2 Medical Device Classification

| Field | Value |
| --- | --- |
| Standard | IEC 62304 |
| Initial Class | Class C — Failing software may result in death of the patient |
| After Mitigation | Class B — Alarming and diagnosis functions mitigated to independent systems |

# 4. Product Context

# 5. System Requirements

| Requirement | Target |
| --- | --- |
| System Activation | < 10 seconds |
| Vital Signs Display | < 1 second after acquisition |
| Vital Sign Alarm Display | < 1 second after detection |
| Connection Alarm Trigger | 5 seconds of inactivity |
| Connection / Misplacement Alarm Display | < 1 second after trigger |
| Diagnose Candidates Display | < 1 second after receipt from AI |
| Diagnose Candidates → HIS | < 1 second |
| DPROC latency budget | ≤ 200 ms |
| DPREC latency budget | ≤ 800 ms |
| DEC latency budget | ≤ 800 ms |
| Diagnosis Convergence (ALGOS) | 2 minutes maximum |
| Vital Signs Input Rate | ≥ 0.1 Hz |
| UI Update Interval | 1 second (timer-based) |

# 6. Constraints and Assumptions

| Type | Statement |
| --- | --- |
| Constraint | First release must use commercially available validated AI model (Open Evidence) — no custom model |
| Constraint | Software must comply with ANSI AAMI IEC 62304 as Class C (mitigated to B) |
| Constraint | Must support 6 device types: ECG, Pulse Oximeter, BP Monitor, Thermal Probe, Capnometer, EEG |
| Constraint | 2-minute diagnosis convergence is an ALGOS budget — not an MMSS budget |
| Constraint | MMSS display latency budget is 1 second (DPREC ≤ 800 ms, DPROC ≤ 200 ms) |
| Assumption | Clinician users are trained medical professionals |
| Assumption | Host system is a compact embedded CPU with real-time OS capabilities |
| Assumption | Open Evidence provides structured diagnose candidates to MMSS |
| Assumption | HIS interface protocol TBD (HL7/FHIR) — ICDs to be defined |
| Assumption | ALGOS sends audible notification independently on 2-minute timeout — no residual MMSS risk |

# 7. Risk Register

| ID | Risk | Prob. | Impact | Mitigation |
| --- | --- | --- | --- | --- |
| R_01 | AI algorithm regulatory clearance delays | Medium | High | Use commercially validated Open Evidence model; engage regulatory early |
| R_02 | AI model validation complexity | High | High | Off-the-shelf validated model mandated for v1.0 |
| R_03 | Connectivity / real-time integration challenges | High | High | Prototype early; prioritise ICD definition for external interfaces |
| R_04 | Sensor misplacement not detected → wrong diagnosis | Medium | Critical | DEVICES must auto-detect misplacement |
| R_05 | Clinician misinterpretation of AI candidates | Medium | Critical | Mandatory simulation training mode in MMSS; hands-on training |
| R_06 | Diagnosis timeout not communicated → delay of treatment | Low | Critical | ALGOS and MMSS both notify on 2-minute timeout |
| R_07 | Connection failure not detected → missed vital signs | Low | Critical | DEVICES generate independent audible alarms; MMSS shows connection alarm |
| R_08 | Portability targets exceed budget | Medium | Medium | Confirm hardware platform weight/size targets early |